This guide is intended to show the user how to set up the CPAP circuit. A separate guide on how to set up the CPAP Dreamstation is available.

The addition of an HMEF bacterial / viral filter between the mask and the exhalation port reduces the risk  of aerosolisation of Covid-19 virus.

Equipment required

| Dreamstation CPAP ventilator |  |
| --- | --- |
| Interface (Mask, full face mask, Hood (if available)) For size guide re mask fit please see mask packaging |  |
| Connecting tubing with bacterial/viral filter |  |
| HMEF Filter (square or round) These are bi-directional (It doesn’t matter which way round they are fitted) Can be used up to 72 hours (if remains dry) |  |

Set up

- Tubing / filter

| Open tubing set | A |
| --- | --- |
| Disconnect narrow tubing from side port at non filter end (A) |  |
| Close / cap off the inlet pipe Discard narrow tubing. |  |
| Important step Connect HMEF filter between mask and tubing. HMEF Filter is bi-directional – It doesn’t matter which way round it is fitted. |  |
| Filtered air from the patient will escape through the side port of the tubing. |  |
| Connect fixed filter end of tubing to CPAP machine |  |

- Entrained Oxygen

| For mask: Attach entrained oxygen from wall directly to mask port |  |
| --- | --- |
| For full face mask: Attach entrained oxygen from wall to side of valve |  |